# FAST

# Mass Spectrometer (TEAMS)

# Known Data Problems and Limitations

### Mismatch Between PAC Voltage and Lookup Table

Cause: PAC discharge, change in nominal PAC voltage, or DPU reset

APID's Affected: All

Period Affected: See the Lockheed TEAMS site http://teams.spasci.com for a list of affected intervals

Comments: Do not use data from these periods

### Self-Stimulation

Cause: A test mode occasionally requested by the experimenters

APID's Affected: All

Comments: Creates an artificial signal at a commandable pixel and time of flight. Should be obvious as the signal overwhelms any naturally occurring signals. Do not use data from these periods

### Changes in Nominal MCP Voltage

Cause: Request by experimenters

APID's Affected: All TEAMS science data

Period Affected: See the Lockheed TEAMS site http://teams.spasci.com for a list of affected intervals

Comments: Do not compare data across an MCP voltage change. Do not use data from 16-26 June 1998 (orbits 7188-7296), as the MCP voltage was changed several times during this period in order to conduct an instrumental experiment.

### H+ Counts Contaminate He++ Channel

Cause: Inherent width of H+ TOF peak

APID's Affected: 1027, 1029, 1047

Period Affected: Entire mission

Comments: Use HiMass data to check degree of contamination; if no separate He++ peak, do not use He++ data. A similar problem affected He+ data from the Freja mass spectrometer

### Elevated Noise Floor at All Times of Flight

Cause: Accidental coincidences at high counting rates

APID's Affected: All

Period Affected: Cusp and radiation belt crossings; periods affected by ram or spacecraft charging

Workaround: Some routines allow background subtraction (check the IDL documentation for details)

### Counts Displaced into Wrong Solid Angle Bin

Cause: Timing error in accumulator board logic

APID's affected: 1027

Period Affected: Intermittently from spring 1997 onward

Workaround: A technique described in K. Seki et al., J. Geophys. Res. 105, 15931, 2000, allows recovery of the initial distribution by assuming the angular distribution is identical to that of the IESA

### Peaks at 6 and 9 AMU/q ("C++ " and "Be+ ")

Cause: Resonance frequencies in TEAMS electronics

APID's Affected: 1028

Period Affected: Intermittent periods during entire mission

### Long TOF Tail for O+

Cause: Energy loss when passing through carbon foil

APID's Affected: 1028

Period Affected: Entire mission

Comments: This limitation has prevented the detection of molecular ions

### Efficiency Drifts

Cause: Aging of MCP's; for anodes nearest spin plane, ram accelerated the aging process. Temperature variations also contribute

APID's Affected: All TEAMS science data

Period Affected: Starting winter 1997 and progressively worsening

Comments: A fix is in progress and will be released as soon as we have completed testing

### Intermittent Dropouts of High-Energy Counts

Cause: Not determined; believed to be timing error in DPU logic

APID's Affected: 1027

Period Affected: Northern hemisphere passes from December 1996 onward

Comments: Believed to be interference from certain fields modes, aggravated during contacts

### Ram

Cause: Spacecraft velocity with respect to background plasma

APID's Affected: 1027, 1028, 1030, 1031, 1047

Period Affected: Traversals of any region with significant cold (< 10eV) O+

Comment: Primarily affects O+ , since the kinetic energy of O+ at 7 km/s (a typical spacecraft speed) is about 8 eV. Can skew moments calculations since these codes assume that the spacecraft velocity is negligible compared to particle velocities

### Spacecraft Charging

Cause: Photoelectron emission and/or problems with E-field sphere bias

APID's Affected: All

Period Affected: Intermittently during fall 1996; all traversals of density cavity when spacecraft is in darkness; all data from fall 2000 onward

Comments: Obvious in the low energy ion spectra. Current moment calculations do not account for the spacecraft potential

### Incorrect Magnetic Field Phase in the Header

Cause: Data collected while magnetometer is off

APID's Affected: 1027, 1028

Period Affected: Low-latitude (equator-ward of 60 degrees invariant) data from about June 2000 onward

Comments: Primarily affects calculations involving pitch angle. Should be fixable in the same way as for the ESA data.